## reports
id | report
1 | Smoke coming from the smouldering power cable of oven OV-550 in Lab 4; unplugged, the cable was scorched, no flame seen.
2 | Trailing extension cable across the walkway in the teaching lab; a student tripped but did not fall or get hurt.
3 | Argon cylinder found standing unchained beside the bench in the teaching lab; chained up.
4 | Centrifuge C-827 in Lab 4 stopped with an imbalance error and its lid lock jammed; samples could not be run that day.
5 | 1.95 L of 37% hydrochloric acid spilled inside the spill tray in Lab 1; absorbed with spill pads.
6 | 300 mL of 37% hydrochloric acid spilled onto the bench from a cracked bottle in the teaching lab; some splashed onto a student's wrist, washed off at the safety shower, slight redness, no further treatment.
7 | Trailing extension cable across the walkway in pilot plant bay A; a student tripped but did not fall or get hurt.
8 | Fixed gas detector in pilot plant bay B alarmed at 16% LEL; a small leak was found at a flange on line L-109 and tightened.
9 | Centrifuge C-180 in pilot plant bay B stopped with an imbalance error and its lid lock jammed; samples could not be run that day.
10 | Drop of 37% hydrochloric acid splashed into a student's eye while pipetting in the analytical lab; eyes rinsed at the eyewash for 15 minutes, no further treatment needed.
11 | Flash fire at the rotary evaporator in Lab 1; the user's forearm was burned and treated at the university clinic by a nurse.
12 | Hot silicone oil spilled from bath B-223 onto a technician's hand in the teaching lab; burn cooled under running water for 20 minutes, first aid only.
13 | Technician's hand caught in the belt guard gap of mixer M-358; three stitches at the clinic and one day off work.
14 | Centrifuge C-275 in Lab 2 stopped with an imbalance error and its lid lock jammed; samples could not be run that day.
15 | Flame at the neck of a flask of acetone on the hot plate in Lab 4; the user switched off the hot plate and covered the neck with a watch glass, flame out at once.
16 | Centrifuge C-118 in pilot plant bay A stopped with an imbalance error and its lid lock jammed; samples could not be run that day.
17 | Centrifuge C-173 in Lab 2 stopped with an imbalance error and its lid lock jammed; samples could not be run that day.
18 | 150 mL of 98% sulfuric acid spilled on the floor of Lab 3 when a bottle was knocked over; cleaned up with the spill kit.
19 | Spare glass condenser found cracked in the cupboard of the analytical lab; thrown away.
20 | 150 mL of 37% hydrochloric acid spilled on the floor of Lab 2 when a bottle was knocked over; cleaned up with the spill kit.
21 | Drum of 37% hydrochloric acid punctured by a forklift in the store yard; about 22 L ran onto the ground before the leak was plugged.
22 | A cardboard box caught fire next to a Bunsen burner in Lab 4; put out with a CO2 extinguisher within seconds, no damage.
23 | Steam hissing from a valve gland on the jacket of reactor R-375 in pilot plant bay A; gland tightened.
24 | Unlabelled bottle of clear liquid found on a shelf in Lab 2; set aside for disposal.
25 | Flame at the neck of a flask of methanol on the hot plate in the teaching lab; the user switched off the hot plate and covered the neck with a watch glass, flame out at once.
26 | Forklift dented the side panel of storage rack 4 in the store; rack still sound, nothing fell.
27 | Relief valve PSV-442 on reactor R-224 did not lift during its bench test; removed and replaced.
28 | Weep of ethanol from the drain valve of tank T-164; about 250 mL collected in the drip tray underneath.
29 | Weep of acetone from the drain valve of tank T-556; about 50 mL collected in the drip tray underneath.
30 | Drop of aqueous ammonia splashed into a student's eye while pipetting in Lab 4; eyes rinsed at the eyewash for 15 minutes, no further treatment needed.
31 | Spare glass condenser found cracked in the cupboard of Lab 3; thrown away.
32 | Paper towels caught fire next to a Bunsen burner in Lab 3; put out with a CO2 extinguisher within seconds, no damage.
33 | Contractor started grinding near the solvent store without a hot-work permit; work stopped by the area owner, no sparks reached the store.
34 | Carbon dioxide cylinder found standing unchained beside the bench in pilot plant bay A; chained up.
35 | Operator strained his back lifting a 25 kg drum in Lab 4; first aider applied a cold pack, he finished the shift.
36 | 750 mL of acetone spilled on the bench in Lab 3 and was ignited by a nearby heater; put out with a dry powder extinguisher.
37 | Feed pump P-747 tripped on motor overload in pilot plant bay A; the run was stopped for 2 hours.
38 | Fume hood FH-469 in Lab 4: sash fully open and face velocity well below the limit, but the low-flow alarm did not sound.
39 | 700 mL of hexane spilled on the bench in the teaching lab and was ignited by a nearby heater; put out with a dry powder extinguisher.
40 | Operator strained his back lifting a 25 kg drum in Lab 2; first aider applied a cold pack, he finished the shift.
41 | Weep of methanol from the drain valve of tank T-735; about 0.05 L collected in the drip tray underneath.
42 | Centrifuge C-866 in Lab 3 stopped with an imbalance error and its lid lock jammed; samples could not be run that day.
43 | Burning smell from hot plate HP-374 in pilot plant bay B; no smoke or flame, controller found faulty and the experiment had to stop until it was replaced.
44 | Steam hissing from a valve gland on the jacket of reactor R-295 in pilot plant bay A; gland tightened.
45 | Operator strained his back lifting a 25 kg drum in pilot plant bay B; first aider applied a cold pack, he finished the shift.
46 | Relief valve PSV-768 on reactor R-497 did not lift during its bench test; removed and replaced.
47 | Forklift dented the side panel of storage rack 1 in the store; rack still sound, nothing fell.
48 | Steam hissing from a valve gland on the jacket of reactor R-876 in pilot plant bay B; gland tightened.
49 | Student slipped on a wet floor in pilot plant bay A and was taken to hospital; X-ray showed a broken wrist.
50 | Burning smell from hot plate HP-702 in pilot plant bay B; no smoke or flame, controller found faulty and the experiment had to stop until it was replaced.
51 | Relief valve PSV-601 on reactor R-891 did not lift during its bench test; removed and replaced.
52 | Steam hissing from a valve gland on the jacket of reactor R-315 in pilot plant bay B; gland tightened.
53 | Centrifuge C-774 in Lab 3 stopped with an imbalance error and its lid lock jammed; samples could not be run that day.
54 | Feed pump P-348 tripped on motor overload in pilot plant bay A; the run was stopped for 2 hours.
55 | Argon cylinder found standing unchained beside the bench in pilot plant bay A; chained up.
56 | Sprinkler went off in the teaching lab after a waste bin fire; the building was evacuated for 30 minutes.
57 | Student seen decanting 37% hydrochloric acid in the teaching lab without the required face shield; stopped and reminded, nothing spilled.
58 | Steam hissing from a valve gland on the jacket of reactor R-190 in pilot plant bay A; gland tightened.
59 | Flame at the neck of a flask of ethanol on the hot plate in Lab 2; the user switched off the hot plate and covered the neck with a watch glass, flame out at once.
60 | Trailing extension cable across the walkway in Lab 3; a student tripped but did not fall or get hurt.
61 | Feed pump P-555 tripped on motor overload in pilot plant bay B; the run was stopped for 2 hours.
62 | Flame at the neck of a flask of methanol on the hot plate in Lab 3; the user switched off the hot plate and covered the neck with a watch glass, flame out at once.
63 | Flame at the neck of a flask of methanol on the hot plate in Lab 1; the user switched off the hot plate and covered the neck with a watch glass, flame out at once.
64 | Student felt dizzy after weighing methanol on the open bench in pilot plant bay A; recovered after 10 minutes in fresh air, no treatment.
65 | About 7.85 L of demineralised water leaked from a split hose onto the floor of Lab 2 and ran into the floor drain.
66 | Spare glass condenser found cracked in the cupboard of the teaching lab; thrown away.
67 | 2.15 L of 37% hydrochloric acid spilled inside the spill tray in the teaching lab; absorbed with spill pads.
68 | 400 mL of aqueous ammonia was poured down the lab sink in pilot plant bay A instead of into the waste container.
69 | Sprinkler went off in Lab 1 after a paper waste fire; the building was evacuated for 30 minutes.
70 | Technician's hand caught in the belt guard gap of mixer M-378; three stitches at the clinic and one day off work.
71 | Sprinkler went off in the teaching lab after a paper waste fire; the building was evacuated for 30 minutes.
72 | Technician's hand caught in the belt guard gap of mixer M-633; three stitches at the clinic and one day off work.
73 | Weekly test: the emergency shower in Lab 3 gave only a weak trickle, far below the required flow; tagged out of use until repaired.
74 | Centrifuge C-515 in Lab 3 stopped with an imbalance error and its lid lock jammed; samples could not be run that day.
75 | Technician cut a finger on a broken beaker while washing up in Lab 3; cleaned and covered with a plaster by the first aider.
76 | Technician's hand caught in the belt guard gap of mixer M-217; three stitches at the clinic and one day off work.
77 | Fire in the motor of stirrer ST-574 in pilot plant bay B; the area was evacuated and the fire brigade attended.
78 | Student felt dizzy after weighing acetone on the open bench in Lab 2; recovered after 10 minutes in fresh air, no treatment.
79 | Weekly test: the eyewash station in pilot plant bay A gave only a weak trickle, far below the required flow; tagged out of use until repaired.
80 | Fixed gas detector in pilot plant bay A alarmed at 25% LEL; a small leak was found at a flange on line L-107 and tightened.
81 | Operator strained his back lifting a 25 kg drum in Lab 3; first aider applied a cold pack, he finished the shift.
82 | Drop of aqueous ammonia splashed into a student's eye while pipetting in pilot plant bay B; eyes rinsed at the eyewash for 15 minutes, no further treatment needed.
83 | Student slipped on a wet floor in Lab 3 and was taken to hospital; X-ray showed a broken wrist.
84 | Technician's hand caught in the belt guard gap of mixer M-467; three stitches at the clinic and one day off work.
85 | Drop of 37% hydrochloric acid splashed into a student's eye while pipetting in Lab 1; eyes rinsed at the eyewash for 15 minutes, no further treatment needed.
86 | Spare glass condenser found cracked in the cupboard of pilot plant bay A; thrown away.
87 | Weep of 30% sodium hydroxide from the drain valve of tank T-689; about 350 mL collected in the drip tray underneath.
88 | Fume hood FH-202 in Lab 4: sash fully open and face velocity well below the limit, but the low-flow alarm did not sound.
89 | Fume hood FH-192 in pilot plant bay B: sash fully open and face velocity well below the limit, but the low-flow alarm did not sound.
90 | Feed pump P-776 tripped on motor overload in pilot plant bay B; the run was stopped for 2 hours.
91 | Drum of methanol punctured by a forklift in the store yard; about 22 L ran onto the ground before the leak was plugged.
92 | Student seen decanting 98% sulfuric acid in the analytical lab without the required face shield; stopped and reminded, nothing spilled.
93 | Spare glass condenser found cracked in the cupboard of pilot plant bay B; thrown away.
94 | Student seen decanting 37% hydrochloric acid in Lab 3 without the required face shield; stopped and reminded, nothing spilled.
95 | Sprinkler went off in the analytical lab after a paper waste fire; the building was evacuated for 30 minutes.
96 | During a walk-round a jumper wire was found bypassing the door interlock of oven OV-583 in Lab 4; removed, no one admitted fitting it.
97 | Compressed air hissing from a push-fit connector on the air line in Lab 2; connector replaced.
98 | Argon cylinder found standing unchained beside the bench in Lab 3; chained up.
99 | Steam hissing from a valve gland on the jacket of reactor R-297 in pilot plant bay A; gland tightened.
100 | Student felt dizzy after weighing hexane on the open bench in the teaching lab; recovered after 10 minutes in fresh air, no treatment.
101 | 0.4 L of acetone spilled on the bench in pilot plant bay B and was ignited by a nearby heater; put out with a dry powder extinguisher.
102 | Fire in the motor of stirrer ST-526 in pilot plant bay A; the area was evacuated and the fire brigade attended.
103 | About 3.6 L of demineralised water leaked from a split hose onto the floor of pilot plant bay B and ran into the floor drain.
104 | 250 mL of hexane spilled on the bench in the analytical lab and was ignited by a nearby heater; put out with a dry powder extinguisher.
105 | 600 mL of 37% hydrochloric acid spilled onto the bench from a cracked bottle in the teaching lab; some splashed onto a student's wrist, washed off at the safety shower, slight redness, no further treatment.
106 | Trailing extension cable across the walkway in Lab 1; a student tripped but did not fall or get hurt.
107 | Fitter began opening pump P-328 without isolating and locking out the motor; supervisor stopped the job before any part was removed.
108 | Feed pump P-884 tripped on motor overload in pilot plant bay B; the run was stopped for 2 hours.
109 | Operator strained his back lifting a 25 kg drum in pilot plant bay A; first aider applied a cold pack, he finished the shift.
110 | Weekly test: the emergency shower in Lab 1 gave only a weak trickle, far below the required flow; tagged out of use until repaired.
111 | Technician's hand caught in the belt guard gap of mixer M-444; three stitches at the clinic and one day off work.
112 | Weekly test: the eyewash station in Lab 3 gave only a weak trickle, far below the required flow; tagged out of use until repaired.
113 | Reactor R-699 reached 6 °C above its high-temperature trip setting, but the trip did not act; operator shut the heater off by hand.
114 | Fitter began opening pump P-437 without isolating and locking out the motor; supervisor stopped the job before any part was removed.
115 | 150 mL of 37% hydrochloric acid was poured down the lab sink in Lab 1 instead of into the waste container.
116 | Fixed gas detector in pilot plant bay B alarmed at 29% LEL; a small leak was found at a flange on line L-823 and tightened.
117 | Fire in the motor of stirrer ST-643 in pilot plant bay B; the area was evacuated and the fire brigade attended.
118 | Flame at the neck of a flask of acetone on the hot plate in Lab 1; the user switched off the hot plate and covered the neck with a watch glass, flame out at once.
119 | Small fire on the hot plate in Lab 4 put out with a CO2 extinguisher; the student had a small burn on one finger treated with burn gel by the first aider.
120 | Small blue flame seen at the burner hose joint in Lab 3; it went out when the gas tap was closed, hose replaced.
121 | Technician's hand caught in the belt guard gap of mixer M-291; three stitches at the clinic and one day off work.
122 | Fixed gas detector in pilot plant bay A alarmed at 21% LEL; a small leak was found at a flange on line L-308 and tightened.
123 | Technician's hand caught in the belt guard gap of mixer M-546; three stitches at the clinic and one day off work.
124 | Centrifuge C-176 in Lab 3 stopped with an imbalance error and its lid lock jammed; samples could not be run that day.
125 | Trailing extension cable across the walkway in Lab 4; a student tripped but did not fall or get hurt.
126 | Centrifuge C-160 in Lab 4 stopped with an imbalance error and its lid lock jammed; samples could not be run that day.
127 | Drum of 30% sodium hydroxide punctured by a forklift in the store yard; about 36 L ran onto the ground before the leak was plugged.
128 | About 3.9 L of 5% salt solution leaked from a split hose onto the floor of Lab 4 and ran into the floor drain.
129 | Steam hissing from a valve gland on the jacket of reactor R-174 in pilot plant bay B; gland tightened.
130 | Drum of 98% sulfuric acid punctured by a forklift in the store yard; about 36 L ran onto the ground before the leak was plugged.
131 | Fire in the motor of stirrer ST-202 in pilot plant bay A; the area was evacuated and the fire brigade attended.
132 | Compressed air hissing from a push-fit connector on the air line in pilot plant bay A; connector replaced.
133 | Weep of hexane from the drain valve of tank T-352; about 300 mL collected in the drip tray underneath.
134 | A clipboard fell from the mezzanine of pilot plant bay A into the barricaded area below; nobody was there.
135 | Technician cut a finger on a broken beaker while washing up in the analytical lab; cleaned and covered with a plaster by the first aider.
136 | Operator strained his back lifting a 25 kg drum in the analytical lab; first aider applied a cold pack, he finished the shift.
137 | Technician's hand caught in the belt guard gap of mixer M-178; three stitches at the clinic and one day off work.
138 | Small fire on the hot plate in the analytical lab put out with a CO2 extinguisher; the student had a small burn on one finger treated with burn gel by the first aider.
139 | Unlabelled flask of clear liquid found on a shelf in the analytical lab; set aside for disposal.
140 | Reactor R-340 reached 15 °C above its high-temperature trip setting, but the trip did not act; operator shut the heater off by hand.
141 | Fitter began opening pump P-853 without isolating and locking out the motor; supervisor stopped the job before any part was removed.
142 | Weep of 37% hydrochloric acid from the drain valve of tank T-340; about 100 mL collected in the drip tray underneath.
143 | Student slipped on a wet floor in the teaching lab and was taken to hospital; X-ray showed a broken wrist.
144 | Weep of acetone from the drain valve of tank T-382; about 50 mL collected in the drip tray underneath.
145 | Fume hood FH-821 in Lab 2: sash fully open and face velocity well below the limit, but the low-flow alarm did not sound.
146 | Argon cylinder found standing unchained beside the bench in pilot plant bay B; chained up.
147 | Fume hood FH-642 in the teaching lab: sash fully open and face velocity well below the limit, but the low-flow alarm did not sound.
148 | Fire in the motor of stirrer ST-742 in pilot plant bay A; the area was evacuated and the fire brigade attended.
149 | Hot silicone oil spilled from bath B-367 onto a technician's hand in pilot plant bay A; burn cooled under running water for 20 minutes, first aid only.
150 | Hot silicone oil spilled from bath B-532 onto a technician's hand in Lab 2; burn cooled under running water for 20 minutes, first aid only.
151 | 500 mL of ethanol spilled on the floor of Lab 2 when a bottle was knocked over; cleaned up with the spill kit.
152 | During a walk-round a jumper wire was found bypassing the door interlock of oven OV-133 in pilot plant bay B; removed, no one admitted fitting it.
153 | Drum of ethanol punctured by a forklift in the store yard; about 36 L ran onto the ground before the leak was plugged.
154 | Hot silicone oil spilled from bath B-264 onto a technician's hand in Lab 2; burn cooled under running water for 20 minutes, first aid only.
155 | Student felt dizzy after weighing acetone on the open bench in pilot plant bay B; recovered after 10 minutes in fresh air, no treatment.
156 | Fixed gas detector in pilot plant bay B alarmed at 17% LEL; a small leak was found at a flange on line L-307 and tightened.
157 | Trailing extension cable across the walkway in pilot plant bay B; a student tripped but did not fall or get hurt.
158 | Sprinkler went off in Lab 3 after a waste bin fire; the building was evacuated for 30 minutes.
159 | Centrifuge C-572 in the analytical lab stopped with an imbalance error and its lid lock jammed; samples could not be run that day.
160 | Drum of methanol punctured by a forklift in the store yard; about 42 L ran onto the ground before the leak was plugged.
161 | Argon cylinder found standing unchained beside the bench in the analytical lab; chained up.
162 | 400 mL of 37% hydrochloric acid was poured down the lab sink in Lab 3 instead of into the waste container.
163 | Weep of 37% hydrochloric acid from the drain valve of tank T-349; about 350 mL collected in the drip tray underneath.
164 | Feed pump P-107 tripped on motor overload in pilot plant bay B; the run was stopped for 2 hours.
165 | Relief valve PSV-650 on reactor R-659 did not lift during its bench test; removed and replaced.
166 | Student slipped on a wet floor in Lab 2 and was taken to hospital; X-ray showed a broken wrist.
167 | 450 mL of aqueous ammonia was poured down the lab sink in Lab 2 instead of into the waste container.
168 | Burning smell from hot plate HP-146 in Lab 4; no smoke or flame, controller found faulty and the experiment had to stop until it was replaced.
169 | Unlabelled bottle of yellow liquid found on a shelf in Lab 3; set aside for disposal.
170 | Sprinkler went off in Lab 4 after a cardboard box fire; the building was evacuated for 30 minutes.
171 | Flash fire at the rotary evaporator in pilot plant bay B; the user's forearm was burned and treated at the university clinic by a nurse.
172 | Student seen decanting 30% sodium hydroxide in Lab 2 without the required face shield; stopped and reminded, nothing spilled.
173 | Flame at the neck of a flask of ethanol on the hot plate in pilot plant bay B; the user switched off the hot plate and covered the neck with a watch glass, flame out at once.
174 | Compressed air hissing from a push-fit connector on the air line in the analytical lab; connector replaced.
175 | Steam hissing from a valve gland on the jacket of reactor R-381 in pilot plant bay A; gland tightened.
176 | Forklift dented the side panel of storage rack 6 in the store; rack still sound, nothing fell.
177 | Drop of 30% sodium hydroxide splashed into a student's eye while pipetting in Lab 1; eyes rinsed at the eyewash for 15 minutes, no further treatment needed.
178 | Flame at the neck of a flask of toluene on the hot plate in Lab 4; the user switched off the hot plate and covered the neck with a watch glass, flame out at once.
179 | 550 mL of toluene spilled on the bench in pilot plant bay B and was ignited by a nearby heater; put out with a dry powder extinguisher.
180 | Fire in the motor of stirrer ST-460 in pilot plant bay A; the area was evacuated and the fire brigade attended.
181 | Sprinkler went off in the teaching lab after a cardboard box fire; the building was evacuated for 30 minutes.
182 | Cooling water pump P-141 tripped; the standby pump started automatically and the experiment carried on without interruption.
183 | Centrifuge C-761 in Lab 4 stopped with an imbalance error and its lid lock jammed; samples could not be run that day.
184 | Unlabelled jar of clear liquid found on a shelf in Lab 2; set aside for disposal.
185 | Hot silicone oil spilled from bath B-125 onto a technician's hand in Lab 1; burn cooled under running water for 20 minutes, first aid only.
186 | Cooling water pump P-274 tripped; the standby pump started automatically and the experiment carried on without interruption.
187 | Fitter began opening pump P-298 without isolating and locking out the motor; supervisor stopped the job before any part was removed.
188 | 450 mL of 98% sulfuric acid spilled inside the fume hood in Lab 1; absorbed with spill pads.
189 | Flash fire at the rotary evaporator in the analytical lab; the user's forearm was burned and treated at the university clinic by a nurse.
190 | Feed pump P-308 tripped on motor overload in pilot plant bay A; the run was stopped for 2 hours.
191 | Flash fire at the rotary evaporator in Lab 4; the user's forearm was burned and treated at the university clinic by a nurse.
192 | 450 mL of methanol spilled on the floor of Lab 4 when a bottle was knocked over; cleaned up with the spill kit.
193 | About 3.45 L of demineralised water leaked from a split hose onto the floor of pilot plant bay A and ran into the floor drain.
194 | Cooling water pump P-257 tripped; the standby pump started automatically and the experiment carried on without interruption.
195 | 1950 mL of 37% hydrochloric acid spilled inside the spill tray in Lab 1; absorbed with spill pads.
196 | 300 mL of acetone spilled on the floor of Lab 2 when a bottle was knocked over; cleaned up with the spill kit.
197 | Fire in the motor of stirrer ST-582 in pilot plant bay A; the area was evacuated and the fire brigade attended.
198 | Small blue flame seen at the burner hose joint in the teaching lab; it went out when the gas tap was closed, hose replaced.
199 | Relief valve PSV-703 on reactor R-866 did not lift during its bench test; removed and replaced.
200 | Weep of 37% hydrochloric acid from the drain valve of tank T-501; about 300 mL collected in the drip tray underneath.
201 | Small fire on the hot plate in pilot plant bay A put out with a CO2 extinguisher; the student had a small burn on one finger treated with burn gel by the first aider.
202 | Compressed air hissing from a push-fit connector on the air line in Lab 1; connector replaced.
203 | Fixed gas detector in pilot plant bay B alarmed at 23% LEL; a small leak was found at a flange on line L-271 and tightened.
204 | Fixed gas detector in pilot plant bay A alarmed at 25% LEL; a small leak was found at a flange on line L-285 and tightened.
205 | Compressed air hissing from a push-fit connector on the air line in the teaching lab; connector replaced.
206 | Student felt dizzy after weighing ethanol on the open bench in pilot plant bay A; recovered after 10 minutes in fresh air, no treatment.
207 | Weep of 98% sulfuric acid from the drain valve of tank T-453; about 400 mL collected in the drip tray underneath.
208 | Weep of 30% sodium hydroxide from the drain valve of tank T-444; about 100 mL collected in the drip tray underneath.
209 | Student felt dizzy after weighing acetone on the open bench in Lab 3; recovered after 10 minutes in fresh air, no treatment.
210 | 1.05 L of 98% sulfuric acid spilled inside the spill tray in Lab 4; absorbed with spill pads.
211 | Fire in the motor of stirrer ST-871 in pilot plant bay A; the area was evacuated and the fire brigade attended.
212 | Feed pump P-301 tripped on motor overload in pilot plant bay B; the run was stopped for 2 hours.
213 | 400 mL of 30% sodium hydroxide spilled onto the bench from a cracked bottle in the teaching lab; some splashed onto a student's wrist, washed off at the safety shower, slight redness, no further treatment.
214 | A cardboard box caught fire next to a Bunsen burner in Lab 3; put out with a CO2 extinguisher within seconds, no damage.
215 | A clipboard fell from the mezzanine of pilot plant bay B into the barricaded area below; nobody was there.
216 | Carbon dioxide cylinder found standing unchained beside the bench in pilot plant bay B; chained up.
217 | Student felt dizzy after weighing hexane on the open bench in Lab 1; recovered after 10 minutes in fresh air, no treatment.
218 | Centrifuge C-166 in pilot plant bay B stopped with an imbalance error and its lid lock jammed; samples could not be run that day.
219 | Technician cut a finger on a broken beaker while washing up in pilot plant bay B; cleaned and covered with a plaster by the first aider.
220 | 1.05 L of hexane spilled inside the fume hood in Lab 1; absorbed with spill pads.
221 | Fitter began opening pump P-398 without isolating and locking out the motor; supervisor stopped the job before any part was removed.
222 | Cooling water pump P-775 tripped; the standby pump started automatically and the experiment carried on without interruption.
223 | 200 mL of 98% sulfuric acid spilled on the floor of the analytical lab when a bottle was knocked over; cleaned up with the spill kit.
224 | Weep of 37% hydrochloric acid from the drain valve of tank T-525; about 0.35 L collected in the drip tray underneath.
225 | Unlabelled bottle of cloudy liquid found on a shelf in Lab 2; set aside for disposal.
226 | Forklift dented the side panel of storage rack 7 in the store; rack still sound, nothing fell.
227 | Fixed gas detector in pilot plant bay B alarmed at 19% LEL; a small leak was found at a flange on line L-317 and tightened.
228 | Weekly test: the eyewash station in Lab 2 gave only a weak trickle, far below the required flow; tagged out of use until repaired.
229 | A spanner fell from the mezzanine of pilot plant bay B into the barricaded area below; nobody was there.
230 | A cardboard box caught fire next to a Bunsen burner in the analytical lab; put out with a CO2 extinguisher within seconds, no damage.
231 | Sprinkler went off in Lab 3 after a cardboard box fire; the building was evacuated for 30 minutes.
232 | Nitrogen cylinder found standing unchained beside the bench in the analytical lab; chained up.
233 | Smoke coming from the smouldering power cable of oven OV-712 in the analytical lab; unplugged, the cable was scorched, no flame seen.
234 | Technician's hand caught in the belt guard gap of mixer M-611; three stitches at the clinic and one day off work.
235 | Technician cut a finger on a broken beaker while washing up in Lab 2; cleaned and covered with a plaster by the first aider.
236 | Fire in the motor of stirrer ST-846 in pilot plant bay B; the area was evacuated and the fire brigade attended.
237 | Weep of 98% sulfuric acid from the drain valve of tank T-812; about 300 mL collected in the drip tray underneath.
238 | Smoke coming from the smouldering power cable of oven OV-371 in Lab 3; unplugged, the cable was scorched, no flame seen.
239 | 500 mL of acetone spilled on the bench in Lab 1 and was ignited by a nearby heater; put out with a dry powder extinguisher.
240 | Technician's hand caught in the belt guard gap of mixer M-310; three stitches at the clinic and one day off work.
241 | Fitter began opening pump P-659 without isolating and locking out the motor; supervisor stopped the job before any part was removed.
242 | Steam hissing from a valve gland on the jacket of reactor R-214 in pilot plant bay B; gland tightened.
243 | 750 mL of methanol spilled on the bench in Lab 4 and was ignited by a nearby heater; put out with a dry powder extinguisher.
244 | Nitrogen cylinder found standing unchained beside the bench in Lab 2; chained up.
245 | Burning smell from hot plate HP-293 in Lab 1; no smoke or flame, controller found faulty and the experiment had to stop until it was replaced.
246 | Weep of toluene from the drain valve of tank T-675; about 400 mL collected in the drip tray underneath.
247 | About 7.85 L of cooling water leaked from a split hose onto the floor of pilot plant bay B and ran into the floor drain.
248 | Burning smell from hot plate HP-495 in pilot plant bay A; no smoke or flame, controller found faulty and the experiment had to stop until it was replaced.
249 | 250 mL of aqueous ammonia spilled on the floor of Lab 3 when a bottle was knocked over; cleaned up with the spill kit.
250 | Flame at the neck of a flask of hexane on the hot plate in Lab 1; the user switched off the hot plate and covered the neck with a watch glass, flame out at once.
251 | Reactor R-590 reached 7 °C above its high-temperature trip setting, but the trip did not act; operator shut the heater off by hand.
252 | Spare glass condenser found cracked in the cupboard of Lab 4; thrown away.
253 | 150 mL of aqueous ammonia spilled onto the bench from a cracked bottle in Lab 2; some splashed onto a student's wrist, washed off at the safety shower, slight redness, no further treatment.
254 | Flame at the neck of a flask of toluene on the hot plate in the analytical lab; the user switched off the hot plate and covered the neck with a watch glass, flame out at once.
255 | Student slipped on a wet floor in the analytical lab and was taken to hospital; X-ray showed a broken wrist.
256 | Weekly test: the emergency shower in pilot plant bay B gave only a weak trickle, far below the required flow; tagged out of use until repaired.
257 | Student slipped on a wet floor in Lab 1 and was taken to hospital; X-ray showed a broken wrist.
258 | Nitrogen cylinder found standing unchained beside the bench in pilot plant bay A; chained up.
259 | Technician cut a finger on a broken beaker while washing up in Lab 4; cleaned and covered with a plaster by the first aider.
260 | A torch fell from the mezzanine of pilot plant bay B into the barricaded area below; nobody was there.
261 | Sprinkler went off in Lab 2 after a cardboard box fire; the building was evacuated for 30 minutes.
262 | Technician cut a finger on a broken beaker while washing up in pilot plant bay A; cleaned and covered with a plaster by the first aider.
263 | Forklift dented the side panel of storage rack 2 in the store; rack still sound, nothing fell.
264 | Student felt dizzy after weighing toluene on the open bench in Lab 2; recovered after 10 minutes in fresh air, no treatment.
265 | Hot silicone oil spilled from bath B-548 onto a technician's hand in pilot plant bay A; burn cooled under running water for 20 minutes, first aid only.
266 | Smoke coming from the smouldering power cable of oven OV-520 in the teaching lab; unplugged, the cable was scorched, no flame seen.
267 | Technician cut a finger on a broken beaker while washing up in Lab 1; cleaned and covered with a plaster by the first aider.
268 | Paper towels caught fire next to a Bunsen burner in the teaching lab; put out with a CO2 extinguisher within seconds, no damage.
269 | Student felt dizzy after weighing toluene on the open bench in pilot plant bay B; recovered after 10 minutes in fresh air, no treatment.
270 | Forklift dented the side panel of storage rack 8 in the store; rack still sound, nothing fell.
271 | Reactor R-307 reached 6 °C above its high-temperature trip setting, but the trip did not act; operator shut the heater off by hand.
272 | 500 mL of aqueous ammonia was poured down the lab sink in Lab 4 instead of into the waste container.
273 | Fume hood FH-230 in Lab 1: sash fully open and face velocity well below the limit, but the low-flow alarm did not sound.
274 | About 5.75 L of 5% salt solution leaked from a split hose onto the floor of Lab 2 and ran into the floor drain.
275 | 100 mL of 30% sodium hydroxide spilled on the floor of the teaching lab when a bottle was knocked over; cleaned up with the spill kit.
276 | During a walk-round a jumper wire was found bypassing the door interlock of oven OV-422 in the teaching lab; removed, no one admitted fitting it.
277 | Student felt dizzy after weighing hexane on the open bench in the analytical lab; recovered after 10 minutes in fresh air, no treatment.
278 | Technician's hand caught in the belt guard gap of mixer M-438; three stitches at the clinic and one day off work.
279 | Fixed gas detector in pilot plant bay A alarmed at 17% LEL; a small leak was found at a flange on line L-812 and tightened.
280 | A torch fell from the mezzanine of pilot plant bay A into the barricaded area below; nobody was there.
281 | Burning smell from hot plate HP-477 in Lab 3; no smoke or flame, controller found faulty and the experiment had to stop until it was replaced.
282 | Fitter began opening pump P-594 without isolating and locking out the motor; supervisor stopped the job before any part was removed.
283 | Sprinkler went off in Lab 1 after a cardboard box fire; the building was evacuated for 30 minutes.
284 | Feed pump P-738 tripped on motor overload in pilot plant bay A; the run was stopped for 2 hours.
285 | Forklift dented the side panel of storage rack 5 in the store; rack still sound, nothing fell.
286 | Fume hood FH-832 in pilot plant bay A: sash fully open and face velocity well below the limit, but the low-flow alarm did not sound.
287 | Drum of 98% sulfuric acid punctured by a forklift in the store yard; about 55 L ran onto the ground before the leak was plugged.
288 | Sprinkler went off in pilot plant bay B after a waste bin fire; the building was evacuated for 30 minutes.
289 | Small fire on the hot plate in Lab 1 put out with a CO2 extinguisher; the student had a small burn on one finger treated with burn gel by the first aider.
290 | 200 mL of hexane spilled on the bench in the teaching lab and was ignited by a nearby heater; put out with a dry powder extinguisher.
291 | Burning smell from hot plate HP-545 in Lab 1; no smoke or flame, controller found faulty and the experiment had to stop until it was replaced.
292 | Unlabelled flask of cloudy liquid found on a shelf in Lab 3; set aside for disposal.
293 | Small blue flame seen at the burner hose joint in Lab 4; it went out when the gas tap was closed, hose replaced.
294 | Hot silicone oil spilled from bath B-463 onto a technician's hand in Lab 4; burn cooled under running water for 20 minutes, first aid only.
295 | 1850 mL of 98% sulfuric acid spilled inside the spill tray in Lab 2; absorbed with spill pads.
296 | Drop of aqueous ammonia splashed into a student's eye while pipetting in pilot plant bay A; eyes rinsed at the eyewash for 15 minutes, no further treatment needed.
297 | Helium cylinder found standing unchained beside the bench in Lab 3; chained up.
298 | Technician's hand caught in the belt guard gap of mixer M-783; three stitches at the clinic and one day off work.
299 | 150 mL of 98% sulfuric acid spilled onto the bench from a cracked bottle in the analytical lab; some splashed onto a student's wrist, washed off at the safety shower, slight redness, no further treatment.
300 | Student seen decanting 98% sulfuric acid in Lab 2 without the required face shield; stopped and reminded, nothing spilled.